## Sheet1
Items | Spec | Data
Integral slit uniformity, 11 scans in 26 mm
Annular | ≤ 1.2 | 1.03
Conventional | ≤ 1.2 | 1.14
Illumination intensity, 11 scans in 26 mm
Annular | 2400 | 2716
Conventional | 2400 | 2954
Dose repeatability and accuracy(at dose >= 5 mj/cm2)
Dose repeatability [%] | ≤ 0.7 | 0.11
Dose accuracy[%] | ≤ 2.0 | 0.47
SAMOS | ≤ 5.5 | 3.39
Image plane deviation | ≤ 175 | 105
Astigmatism | ≤ 125 | 54
Best Focus(Dynamic) | 0+/-35nm | 0.017
image sensor measurements
Focus repeatability [nm] | ≤ 30 | 8.06
(Rx) Scanaxis tilt repeatability [μrad] | ≤ 6.0 | 0.2
(Ry) Scan axis tilt repeatability [μrad] | ≤ 2.0 | 0.41
Translation repeatability [nm] | ≤ 10 | 0.21
Magnification repeatability [ppm] | ≤ 1.0 | 0.13
Die rotation repeatability [μrad] | ≤ 1.0 | -0.07
Golden Matching ( R5 )
GM Tx | 0.05 +/- 0.0075 um | 0.0512
GM Ty | 0.05 +/- 0.0075 um | 0.0507
GM Rotation | 0.0 +/- 0.4urad | 0.0605
GM Orthogonality | 0.0 +/- 0.4urad | -0.0081
GM X_EXP | 0.0 +/- 0.4urad | 0.114
GM Y_EXP | 0.0 +/- 0.4urad | -0.078
GM R_ROT | 0.0 +/- 2.0ppm | -0.0323
GM R_MAG | 0.0 +/- 2.0ppm | -0.3292
GM AR_ROT | 0.0 +/- 2.0ppm | 0.0451
GM AR_MAG | 0.0 +/- 2.0ppm | -0.0391
Delta Shift Check / R5 X | 0.0 +/- 100 nm | -23
Delta Shift Check / R5 Y | 0.0 +/- 100 nm | 34
Static XY
Max. DX | 0+/-40nm | 10.5
Max. DY | 0+/-40nm | 11.7
Rotation | 0+/-200urad | 0.321
Max DX-Max Res X | 0+/-15nm | 3.2
Max DY-Max Res Y | 0+/-15nm | 8
Reticle Height | 0+/-5um | -0.265
Skew + Rot | 0+/-0.5urad | -0.09
Scaling + Mag | 0+/-0.5ppm | 0.2
Dynamic XY
Max. DX | 0+/-30nm | 8.8
Max. DY | 0+/-30nm | 9.6
Max DX-Max Res X | 0+/-15nm | 4.1
Max DY-Max Res Y | 0+/-15nm | 2.6
Asym. Rotation | 0+/-0.5urad | 0.1
Asym. Mag. | 0+/-0.5ppm | -0.148
Stone to lens tilt Delta Rx | 0+/-1urad | -0.1
Stone to lens tilt Delta Ry | 0+/-1urad | 0.2
Moving standard deviation
X (mean+3σ) [nm] Y (mean+3σ) [nm] | ≤ 30 ≤ 30 | 16.2
16.3
Moving average |  | .
X (mean+3σ) [nm] Y (mean+3σ) [nm] | ≤ 15 ≤ 15 | 4.9
6.2
`